Tti-1604 communication protocol

# Communication hardware level

The RS232 connector has standard pinout. The interface is opto isolated from the multimeter. The computer has to provide power for the TXD output drive and the RXD input via the DTR and RTS lines.
## Connector pinout
Pin no	Name	Description
1	DCD	Linked to DTR
2	TXD	Transmitted data from instrument
3	RXD	Received data to instrument
4	DTR	DTR must be set to logic 0 from the host (+V)
5	GND	Signal ground
6	DSR	Linked to DTR
7	RTS	RTS must be set to logic 1 from the host (-V)
8	CTS	Linked to RTS
9	GND	Signal ground

## The RS-232 port settings
9600 Baud,1 Start, 1 Stop, 8 Data Bits, Handshake DTR/DSR, Timeout 500ms

In my experience, if USB-Serial adapter is in use the multimeter accept power supply from +-4.5V.

# Commands from host to the instrument
Every button has an own ASCII character. Send any of this character is equal then press that button.
Key	Character	Description
Up	a
Down	b
Auto	c
A	d
mA	e
V	f
Operate	g
Ω	i
Hz	j
Shift	k
AC	l
DC	m
mV	n
Remote mode	u	Activate serial communication
Local mode	v	Deactivate serial communication

Every sent character being echoed back. If no response within 300mSec then command has to resend.

# Data format from instrument
All measurement from the instrument is packed to 10 byte long message.
## Characters in the message
Index	Function
0	Start character
1	Function and Range information
2	Special functions
3	Minus sign
4	Display number 1
5	Display number 2
6	Display number 3
7	Display number 4
8	Display number 5
9	Status information

## Character description
### Start character
Always 0x0D
### Function and Range information
Bit[0..2]: Active functions
Binary value	Function
001	mV
010	V
011	mA
100	A
101	Ω
110	Continuity
111	Diode test

Bit[3]:
1	AC
0	DC

Bit[4..6]: Range information
Binary value	Range
000	400Ω
001	4kΩ / 4VAC / 4VDC / 4mADC / 1mAAC
010	40kΩ / 40VAC / 40VDC / 10ADC / 10AAC
011	400kΩ / 400VAC / 400VDC / 400mADC / 100mAAC / 400mVDC / 400mVAC
100	4MΩ / 750VAC / 1000VDC
101	40MΩ

Bit[7]: Not used
### Special functions
Bit	Function
1	T-Hold
2	Min/Max
4	Hertz
5	Null
6	Auto

### Minus sign
Bit[1]: Minus sign visible
### Display numbers
Every display value is a representation of the seven segment display.

Bit[7]: A
Bit[6]: B
Bit[5]: C
Bit[4]: D
Bit[3]: E
Bit[2]: F
Bit[1]: G
Bit[0]: DP

### Status information
Bit no	Function
0	Double Beep
1	Auto range set
3	Cont Buzz
4	Disp Min
5	Disp Max
6	Disp Hold
7	Gate 10 Sec